Environmental Sample Processor
Contextual Sensors
5/25/22 Brent Roman brent@mbari.org

---
Supported Instruments
Can => internal environmental sensors within ESP core's housing
Temperature, humidity, pressure, battery voltage, amperage
Updates every 10 minutes as long as ESP application runs
CTD => Seabird SBE 16plus V2 interfaced via RS-232 sensor 1
Temperature, pressure, conductivity, plus optional...
Fluorometer, Transmissometer, Oxygen Sensor (1 of 2 types)
ISUS => one of two types interfaced via RS-232 sensor 2
Concentration of nitrate and, optionally, bisulfide
Support for all manufactured at MBARI
Some later models from Satlantic (in use at WHOI)
TBD = Something new can yet be interfaced as RS-232 sensor 3
Note: this port is not currently wired to lid of the can

---
Polling Contextual Sensors
Trickier than it would first seem
ISUS must synchronize with CTD to receive timely updates
Sample rate optionally quickens during sampling
Multiple threads may not access instruments simultaneously
The Can's internal sensor polling is controlled independently
Code is in Polling object in mission/skeleton.rb
Polling.start	#starts SensorPolling with new parameters
Polling.stop	#stops polling and properly closes instrument files
Polling.pause #stops until resumed
Polling.resume #resumes previous polling schedule if paused
Instrument shows last sampled state of all Instruments
CTD, ISUS, Can show last sampled state of each Instrument

---
Internal Environmental sensors
can is short for Gate.queryCan --> forces immediate sampling
can.temperature => internal temp. at top of can in degrees C
can.humidity => humidity in % of saturation
can.pressure => internal pressure in psia
can.voltage => instantaneous battery voltage
can.current => instantaneous battery load in amps
can.avgCurrent => averaged battery load in amps
can.waterAlarm => percent “wet” (0..100) usually < 1
Wattage is merely can.current * can.voltage
Gate.can returns most recent can environment sample taken
Typically updated every 10 minutes
Recorded in binary 'real.log' file
$ dumplog source=:Can #will list Can environment samples
Gate.canPollInterval = desired update rate for Gate.can
Gate.canPollInterval = Delay.new “7:00” #change rate
Set to zero to disable can environmental sampling entirely
This is the default for MFBs lacking can env sensors.

---
Seabird CTD
Seabird 16plus V2 CTD with
support for fluorometer, transmissometer, oxygen sensor, ...
Generates file CTD-*.hex of raw samples
CTD.status # shows instrument status
CTD.pumpmode = mode, where mode is either:
:off, :beforeSample, or :duringSample
s = CTD.sample => returns sample object, assigns it to variable s
s.temperature => sea temperature in degrees C
s.conductivity => conductivity in S/m
s.pressure => pressure in decibars
s.transmissometer => % optical transmission
s.beamAttenuation => extinction coefficient in 1/m
s.sampleTime => time at which this sample was started
s.dataTime => time at which this sample was finished
s.depth => depth in meters (derived from pressure)
s.salinity => salinity in mythical PSUs
More documentation in lib/instrument/ctd.rb

---
ISUS
ISUS = In-Situ Ultraviolet Spectrometer
Stores raw spectra in ISUS-*.dat (MBARI's ISUS only!)
Logs errors in ISUS-*.err
Requires temp., salinity & depth from the CTD !!
ISUS.status # shows instrument status
ISUS.species = 2 (or 3) #three to include bisufide
ISUS.fit = 217..240 #spectral fit window in nm (tweak for species)
ISUS.fromCTD temp, salinity, depth #update ISUS from CTD
s=ISUS.sample => sample with most recent values fromCTD
s.no3 => Nitrate concentration in uM/L
s.br => Bromide in uM/L
s.hs => Bisulfide in uM/L (only valid if species>2 and fit tweaked)
s.sampleTime => when sample was requested
s.dataTime => when sample was recorded
More documentation in lib/instrument/isus.rb

---
Parameters controlling Contextual Sensor Polling
$global variables determine instruments' configuration/polling rates
These may be assigned anytime before Polling.start
But, usually they get set once in mission/phasecfg.rb
Missions with :until=>time automatically invoke Polling.start
CTD
$ctdPumpMode=:duringSample #may be :beforeSample or :off
$ctdInterval=Delay.new “5:00” #sample CTD every 5 minutes
$ctdPeriod=Delay.new “1:00:00” #upload CTD data every hour
$samplingCTDinterval=Delay.new “2:30” #2x faster ...
$samplingCTDperiod=Delay.new “30:00”	# while sampling
ISUS
$isusSpecies = 2 #ignore sulfides by default (3 to include them)
$isusFit =217.240 #because Luke says it should be so :-)
ISUS polling rate is CTD sampling rate + 10 minutes
ISUS auto-sampling cannot be disabled