Physical State
5/24/22 Brent Roman brent@mbari.org

---
ESP Physical State
- Canister environmental sensors
- Power Switches
- All moving actuators:
    - Rotary Valves
    - Solenoid Valves
    - Syringes
    - Clamps
    - Carousel
    - Elevator
    - Elbow
    - Gripper
- Puck Heaters

---
Gateway Canister State
- Quietly logged once every
  - 10min for 2G, 2min for 3G
  -> Gate.queryCan #or simply -> can
  Can@22:40:11, 24.2C, 67% humidity, 14.2psia, 13.684V, 0.256A, 187.834Ah, 3.50W
  - Ah is only available with newest firmware
  -> Gate.queryCan takes an immediate reading
    - can is an alias for Gate.queryCan
  -> Gate.can uses the most recent reading
  -> Gate.canPollInterval=30.seconds
  -> Gate.canPollInterval=0	#no polling

---
Canister State details
-> Can.unit	#all available can sensors
  {avgCurrent: {format: "%.3fA avg"},
  batteryUsed: {format: "%.3fAh"},
  current: {format: "%.3fA"},
  humidity: {format: "%d%% humidity"},
  pressure: {format: "%.1fpsia"},
  temperature: {format: "%.1fC"},
  voltage: {format: "%.3fV"},
  waterAlarm: {format: "%d%% Wet!",
  threshold: 1.5}}
-> Can.missionDuration #duration power has been on
  5 weeks, 3 days, 10:19:32.40625
-> c=can; (c.voltage*c.batteryUsed).round #~Wh used

---
Gateway Power Switches
-> Power	#state of all power switches
  {analytic1: false,
  analytic2: false,
  analytic3: false,
  camera: false,
  core: false,
  raw: false,
  sampler: false}
-> Power.sampler :on #turn sampler on
-> Power.on :sampler #turn sampler on
-> Power.off :sampler, :camera #turn off sampler and camera
-> Power.offSince[:sampler] #when sampler was turned off or nil
-> Power.onSince[:core]	#when core was powered on or nil
-> Power.sampler	#True if sampler is on

---
Linear Actuator Class Hierarchy
  - Slide => Lowest level at which end-users interface with hardware
    - Think of a slide trombone with named positions for arbitrary “notes”
    - Forearm, Elbow, Carousel
  - Clamp => inherits from Slide
    - Adds closed?, open?, and closeAndVerifyPuckPresence
  - Scale => inherits from Slide
      - Adds a linear, numeric scale (in standard units) to Slides
      - Elevator #unit = puck height
  - Syringe => inherits from Scale
      - adds pull, push, fill, empty volume methods
      - Collection, Processing, Sampler, Analytical syringes
  - Thermal => inherits from Scale
      - replaces motor with a heater
  - Errors come from microcontrollers, which are all managed by Slide class
    - This is why Linear Actuators report most errors as “Slide::Error”

---
Other Actuator Classes
- Shaft is a rotary actuator
  - all of which just happen to spin a rotary valves
- Gripper is a two state actuator
  - with binary position sensing
  - May control a robotic hand or a motorized valve
- Solenoid is a two or three state solenoid actuator
  - without position feedback sensing
- Valve is a Solenoid used to control fluid flow
- Valve::Manifold is an array of Valve
  - sharing a common fluid path

---
Using AxisMaps
An AxisMap maps all raw counts to corresponding position names
They are typically accessed via their associated Axis or Positions:
axis.legend => the AxisMap as a Hash
axis.list => list of all names without raw positions
axis.labels => list of only the position labels – omitting aliases
axis.maxPosition => position mapped to greatest raw counts
axis.minPosition => position mapped to least raw counts
axis.advance => move to position with next higher raw counts
axis.retard => move to position with next lower raw counts
axis.at?(position) => true if axis is at (or near) specified position
axis.near?(position) => true if axis is at or near position
axis.between?(pos1,pos2) => true if axis is (nearly) between
axis.rawId(rawCount) => position nearest rawCount (reverse map)
position.advance(detents) => position with next higher raw counts
position.retard(detents) => position with next lower raw counts
position.near?(position) => true if positions very near each other

---
Axis Map Example
- Hardware counts <==> names and aliases
  -> Forearm.legend
  {-12793 => 3, -12782 => 5, -12726 => 7,-12711 => 4,
  -12706 => 2, -12699 => 1, -12678 => 6,
  -12501 => [CC, :collection, Collection],
  -12418 => [PC, :processing, Processing],
  -12342 => [:garage, FlushPuck, FlushPuck::Garage],
  0 => :home,
  2800 => ["retracted", :retract, :clear]}
- Defined in configure.rb as:
  Forearm.detents 0=>:home, 2800=>"retracted",
  -12501=>CC, -12418=>PC, -12342=>:garage,
  -12699=>1, -12706=>2, -12793=>3, -12711=>4, -12782=>5,
  -12678=>6, -12726=>7

---
Solenoid Valves
- Solenoid valve state is either open, closed, or unknown
- Intake and Exhaust external solenoid valves
  - Unipoler, normally closed
  - Draw a lot of power while opened
- Solenoid valves in Manifolds
  - Bipoler, latching
  - Latching Valve state is initially unknown
  - Draw power in short pulses when changing state
  -> Intake.to :open	#holds Intake open
  -> Intake.open	#ditto
  -> Solenoid	#shows Intake and Exhaust state

---
Solenoid Valves Plumbing
- Each ESP 2G Dwarf microcontroller
  - drives 8 solenoids numbered 0..7
- To show how a solenoid is connected:
  -> Intake.wiring #how is Intake wired?
  sampler[6]	#it’s 2nd to last on sampler dwarf

---
Valve::Manifold
- Each composed of a series of Solenoid::Valve plus an endName
  - Valve::Manifold state = either the name
    - of its first opened Solenoid::Valve
    - or its endName, if no Solenoid::Valve is open
- Individual valves in the manifold can be accessed
  -> CSR.series[0].open #open 1st CSR valve
  -> CSR.wiring #shows wiring connections
  [:lysis <=> collection[0],	#these need not be in order
  :diluent <=> collection[1],
  :RNAlater <=> collection[2],
  :mfbkill <=> collection[3],
  :kill <=> collection[4],
  :flush <=> collection[5]]

---
Valve::Manifold Configuration
- Collection Series configuration is often machine specific
- Example for ESPwaldo, defined in its configure.rb file:
  :CSR.denotes Valve::Manifold :Collection, [
  Valve.reagent(:lysis, CollectionValves,0),
  Valve.reagent(:diluent, CollectionValves,1),
  Valve.reagent([:RNAlater, :rnal], CollectionValves,2),
  Valve.reagent(:mfbkill, CollectionValves,3),
  Valve.reagent(:kill, CollectionValves,4),
  Valve.reagent(:flush, CollectionValves,5) ], :air
- :air is the manifold’s “endName” representing
  - its state when all its series valves are closed
- :rnal is an alias for :RNAlater

---
aliases and labels
- Each named position has exactly one label
  - in addition, it may have any number of aliases
- A position is output by its label
  - but may be input by its label or any corresponding alias
  -> CSR.aliases
  {rnal: :RNAlater}
  -> CSR.alias :clean=>:flush #clean now an alias for flush
  -> CSR.relabel :atmosphere, :air
  -> CSR.to :air
  Collection Valve::Manifold selects atmosphere
- rather than air!

---
Rotary Valves
- Each rotary valve is controlled by its Shaft position
  - Shaft is the class of all rotary valves
- Raw shaft state is an angle from 0..511
  - A number of these may be defined as named positions
- Rotation direction to the new goal position
  - may move over raw position 0
  - may be specified
    - to avoid moving over another position
- A goal position may be specified
  - as a named position
  - exactly between 2 named positions
  - as a raw offset from one of the above

---
Shaft Configuration
Each 2G Dwarf may control up to 4 rotary valves numbered 0..3
-> PTV.wiring	#displays wiring information
processing[2]	#wired to 3rd channel of processing dwarf
-> PTV.legend	#displays position map
{64 => [PRV, 1, #positon labeled PRV [with alias 1] is raw angle 64
192 => [:PRVmixing, 2],
320 => [:mixing, 3],
448 => [:puck, 4]}

---
Rotating Shafts
-> PTV.to :mixing	#rotates to mixing position fastest way ‘round
-> PTV.select :mixing, avoiding: PRV
    - #rotates to mixing position
    - in a direction that avoids rotating by the PRV position
-> PTV.select :mixing, via: PRV	#rotates opposite way!
    - #rotates to mixing position via the PRV position
-> PTV.dialBetween :mixing, :puck
    - #rotates to between mixing & puck position fastest way ‘round
-> PTV.dialBetween :mixing, :puck, avoiding: PRV
    - #rotates to between mixing & puck position, avoiding PRV
-> PTV.rawAngle	#return the raw angular position of the shaft
320
-> PTV.at? :mixing	#true if PTV is at, or very near, mixing
true

---
Gripper Characteristics
- Two position actuator with minimal position feedback
- In one of two states
  - or transitioning between those states
- The two states are named when configured
  - they may not have aliases or labels
- Examples are:
  - The ESP 2G Hand was original Gripper actuator
  - Some ESP 2G external Sample valves are controlled as Grippers
  - 3G External rotary valves are all Grippers

---
Gripper Use
-> Hand.close	#closes the Hand Gripper
-> Hand.open	#opens it
-> Hand.state	#:open, :closed, or :unknown
-> Hand.open?	#true if Hand is open
true
- Only on NOAA GLERL 2G ESPs...
  -> Sample.deep	#moves Sample valve to its :deep position
  -> Sample.shallow #moves Sample valve to its :shallow position
- -> Sample.state	#:shallow, :deep, or :unknown
- -> Sample.deep?	#true if Sample valve is in its :deep state
  false

---
Slide Characteristics
- Linear Actuator having precise position feedback
- Raw positions are in hardware counts
  - Each Slide’s count units may be different
  - Most Slides require physical ‘homing’
    - to calibrate their counts position sensor
    - A Slide that is not yet homed is ‘lost’
  - are mapped to names with associated AxisMap
- Each Slide includes a set of named configurations
  - that set motor limits and velocity profile
  - only one such configuration is active at a time

---
Basic Slide Operations
The Slide is the “base class” for linear actuator axes
slide.configure cfg => forces configuration object cfg to dwarf
slide.reconfigure cfg => sends cfg only if changed from last
slide.in(cfg) {block} => execute block in configuration cfg
slide.position => return the slide's current position
slide.goal => return the slide's current goal position
slide.jog counts => move specified # of raw encoder counts
slide.seek goal => move to specified goal position
Without updating servo's configuration
slide.to goal, config => move to specified goal position
Updating servo's configuration if appropriate
slide.hold => hold the current position
slide.coast => turn off the servo
slide.force => apply constant “force” (slide.force 0 = slide .coast)
slide.stop => brake to a stop as fast as possible
slide.log(decimator) {block} => log slide status while doing block
slide.status => return current slide servo status object

---
How Scales Differ from Slides
Scales inherit all the operations of Slide, adding:
Linear mapping of logical “amounts” or “units” to raw counts
rawCount = scale.countsPerUnit * amount + zero
zero is simply the rawCount value at 0 amount
scale.zero => -12580 #example case
scale.gain => scale.countsPerUnit => 32498.0
AxisMap associated with a Scale:
Must contain at least two positions whose labels are numeric
[there should be only two numerically labeled positions]
These positions project the scale's linear mapping onto counts

---
Scale::Skew objects
A Scale::Skew is a generic, linear mapping object
represents
y=mx * b
scale.skew => -12580.000+32498*counts
scale.skew.gain => 32498.0, scale.skew.bias => -12580
scale.skew.apply(2) => 52416.0 # == 2*32498 – 12580
solves for y (engineering units) given x (counts)
scale.skew.reverse(scale.skew.apply(x)) => x
solves for x (counts) given y (engineering units)
Skew.bestFit(counts, units) => skew that best fits data
Skew.interpolate() => interpolates among array of skews
Scale::Skews are also used to calibrate Thermal pads!

---
How Syringes differ from Scales
A syringe is merely a scale with volumetric units
volume is defined as an alias for amount
Similarly for maxVolume and minVolume
fill method moves to the syringe's maxPosition
empty method moves to the syringe's minPosition

---
2G ESP Dwarf DC Motor Servos
- Two identical servo channels
- 64hz sampling timebase (sample rate typically 32hz)
- Each Channel's Inputs:
  - Quadrature incremental encoder
    - (A and B 90 degrees out of phase)
  - Home flag (typically a hall effect sensor)
  - Optional threshold sensor
  - Forward and Reverse limit switches
  - One General Purpose digital input bit (for gripper)
Each Channel Outputs:
  - PWM -100% to 100% (15 kHz with 1% resolution)
  - One General Purpose digital output bit

---
Configuration Object Details
:samplePeriod = number of 64hz timebase tics per sample tic
  - Default value = 2 (Typically 1 or 2)
:encoder, :threshold, :home sensor power / polarity
  - Default value = :off (may be :positive or :negative)
:homeDirection = :forward or :reverse
  - Default value = :reverse
  - :reverse moves negative if home flag inactive
:brake = short motor terminals on servo error (:false or :true)
  - Default value = true
:debug = output servo state at sample rate while seeking goal
  - Default value = false

---
Control Gains and Factors
PID :gain struct with members P, I, and D
  - Default values for each are 0
  - Servo will not operate until at least one is non-zero
  - Effective value of P and D is divided by 4096
  - I is effectively divided by 16384
:friction compensation gain
  - cmdVel * friction / 4096 added to PWM output
  - cmdVel = Commanded velocity
:stiction compensation factor
  - If negative cmdVel, subtract stiction/2 from PWM
  - If positive cmdVel, add stiction/2 to PWM

---
Trajectory Generator (1 of 2)
:acceleration & :deceleration in counts/tic/tic
  - Default values for each are 0, normally positive
  - Specify negative acceleration to disable “softstart”
  - Zero :deceleration implies deceleration=abs(acceleration)
:maxSpeed = plateau velocity in counts/tic
  - Temporarily reduced when PWM limits reached to prevent trajectory errors due to low battery voltage
:minSpeed = slowest acceptable progress rate (counts/sec)
  - Speed error if maxSpeed reduced below minSpeed
:maxSettling = max tics to allow to servo to settle at goal
  - Default 0, typically 2 – 3 seconds worth of tics
  - Just ensures that position error not returned too early

---
Trajectory Generator (2 of 2)
:stopWindow detemines how nearly goal should be reached
  - Specified in encoder counts (16 bit limit max)
  - Temporarily increased each time goal is passed
  - Special Value false indicates no (more) reseeks allowed
  - Defaults to Special Value :deceleration = deceleration rate
  - Also accepts value :acceleration
:hunt determines whether to adjust setpoint after goal reached
  - Defaults to false, set true to “fight” to hold exact position at goal
  - Setpoint is never adjusted if position within stopWindow
:thresholdOffset determines how far from threshold to stop when reached
  - Defaults to 0 encoder counts
  - When threshold reached before goal, goal = position + thresholdOffset
  - Used to position top of puck stack with respect to ESP's top plate

---
Core Limits
:maxPWM & :minPWM
  - Max must be >= min, but each may be negative or positive
  - Constrains servo output, but does not constrain “force” command
  - Effective maxSpeed is reduced when servo reaches these PWM limits
:maxPositionErr determines absolute maximum tolerable servo error in different contexts:
  - SeekErr if stopWindow grows too large due to repeatedly missing goal
  - TrajectoryErr if position becomes too far from setpoint while transiting
  - PositionErr if position moves too far from goal after arrival
:maxCurrent determines maximum allowable motor current
  - In milliamps
  - Should never be set > 2000mA

---
Pressure Limits
:maxInPress, :maxOutPress, :minInPress, :minOutPress
  - 0 to 4095 ADC counts
  - Maximum/Minimum tolerated Intake and Outlet pressures
  - Constraint disabled if corresponding max == min
  - All default to 0
:maxDeltaPress & :minDeltaPress -- (-4095 to 4095)ADC counts
  - Maximum/Minimum tolerated pressure difference
  - Constraint disabled if set to special value: false
  - All default to false (there is no corresponding value true)
Generic “Pressure Error” results if any of the above are violated
  - One must check status to determine the exact problem

---
Pressure Servo Configuration
:inputDeltaPress determines if pressure delta is sensed or derived
  - True to input the difference from ADC 7
  - False to derive it as (intake – outlet) pressure
  - Defaults to false
:pressBias is subtracted from delta pressure before use
  - In servo or limit check
  - Defaults to 0
:pressGain is the proportional gain of the pressure servo
  - Scaled like P and D, pressGain is *4096
  - Reduces acceleration from that normally determined by the trajectory generator.
  - Never causes command velocity to fall below minSpeed

---
Default Processing Syringe I2C::Servo::Configuration
:PS.denotes Syringe "Processing Syringe",
      :processing, 0,
:encoder=>:negative, :home=>:negative, homeDirection:false,
maxPositionErr:65,
gain:PIDgain(3500, 3000, 1300),
friction:170,
maxSpeed:100, minSpeed:30,
acceleration:5,
maxCurrent:120,
maxSettling:3*32
PS.maxDuration=160	#how long can a move take?
- excerpted from shallow/preconfig.rb

---
Alternative Processing Syringe Servo Configurations
- Based on default configuration shown on previous slide:
  PS.defCfg :fast, maxSpeed:300
  PS.defCfg :slow, maxSpeed:50
  PS.defCfg :slow1, maxSpeed:10, minSpeed:2, acceleration:2
- Based on :slow1 configuration (defined above):
  PS.defCfg :slow2, :like=>:slow1, minSpeed:5
- excerpted from shallow/postconfig.rb

---
Switching Servo Configurations
### The easy (and correct way) ###
PS.to PS.maxVolume/2, :slow1 #half full (or is it empty?)
  - Only changes the configuration if necessary
  - Don't use .seek unless sure the config already loaded on dwarf.
### The hard (and also correct way) ###
  PS.in :slow1 do
    PS.to PS.maxVolume/2
    PS.empty #this is still in PSslow1
  end
  PS.fill #old configuration restored (likely PSconfig)
  - slide.in {block} constructs may be nested arbitrarily deep

---
I2C::Servo::Status Objects
- :enabled = true if servo control is active
- :pastRLS, :pastFLS, :pastThreshold, :home
  - True if corresponding switch is closed
- :position = 32-bit signed offset from home position
- :velocity = 16-bit signed in encoder counts/tics
- :current = signed milliamps
  - Always agrees with sign of PWM status below
- :PWM = signed percent PWM duty cycle
- :err = 16-bit signed (setpoint – position)
- :voltage = raw motor voltage (in volts)
  - This is the only floating point value

---
Servo Pressure Status
- Recall that pressure may be a proxy for any arbitrary volage input
- :inPress = intake pressure in raw ADC counts (0-4095)
- :outPress = outlet pressure in ADC counts
- :deltaPress = delta pressure in ADC counts
  - This is always ADC channel 7
  - It is not affected by the :inputDeltaPress configuration flag

---
Plotting Slide Servo Trajectories
- Add ssh key to workstation's authorized_keys file
  - So that the ESP host can run commands without password prompts
    - Test from Linux shell prompt, on ESP host, by invoking:
      $ ssh workstation ls
    - This is a security breach. Remove key when done if it worries you.
- Edit remotePlot method utils/plot.rb as necessary
    - To change the workstation name (and possibly the display number)
  - -> require 'plot' #only once per session
- To produce each new plot window:
  -> remotePlot slide.log {blockOfCodeExercisingSlide}
- e.g. plotting default status fields of position, velocity and current:
  -> plot	CC.log {CC.to :closed}
- e.g. plotting :current,:voltage, :pwm, and :err
  -> plot(CC.log {CC.to :closed}, :current, :voltage, :pwm, :err)