ESP Puck Tracking
4/12/24 Brent Roman brent@mbari.org

---
Why is Puck Tracking important?
- There are many types 2G ESP of pucks
- Pucks must be processed in prescribed sequence
- Dropping a puck ends a deployment
  - typically jambs Elbow or Forearm
- Can flood the can if dropped puck not detected
  - and clamp closed without a puck loaded in it
- We want to detect puck errors early to avoid this!

---
Why is puck tracking difficult?
- ESP cannot distinguish among puck types.
  - cannot read etched serial #, etc.
  - Even the flush puck appears identical
- No direct sensors for puck presence
  - ESP can’t tell if Hand is holding a puck
  - Opening Hand risks dropping it
- ESP power may be remove at any time
  - even while handling pucks :-(

---
Sensing Puck Presence
- Motor current when closing Clamp
  - increases as spring compresses
    - only if there’s a puck in the clamp!
    - anything puck in clamp sized would do
- IR beam under camera
  - blocked by top puck in stack
    - Elevator position at that instant allows ESP to calculate height of puck stack.
    - Assuming pucks are of standard height.

---
Counting Pucks
- Expected puck stack height in a tube
  - updated when puck is added or removed
- Every time puck stack lifts, its height
  - is compared with its expected height
  - Puck::Error raised if height != expected
    - within 0.5 puck height units
- This same error will occur if an operator
  - manually changes a tube’s stack height
  - unless the system is informed

---
Tracking Pucks in play
- Pucks currently in play are tracked from
  - tube stack or garage to Hand to destination Clamp
  - unclamping, back to Hand, to tube stack or garage
- Each puck move is recorded in $ESPmode.slot file
- When ESPserver starts,
  - it replays the .slot file to determine location of all pucks
  - If any are out of place (i.e. in play)
    - ESP.ready! will avoid dropping pucks
      - reclamps previously clamped pucks
      - Hand refuses to open if it previously held a puck.

---
Resetting Puck Tracking
-> Puck.count	#counts all pucks in carousel
-> Puck.count 2	#count only pucks in tube 2
-> access 2	#present tube 2 and forget its height
-> access	#present currently selected tube, forget its height
- These do not move any actuator:
  -> clear!	#forget all puck heights
  -> clear! 2..4	#forget puck height of tubes 2, 3 and 4
- If you see any of these sorts of messages on startup:
  Storage::Warning in simfast – (clamp|Hand).holds *puck
  - the system thinks that pucks were in play when it last exited
  - If you have already manually corrected this situation,
  -> Storage.recovery.clear	#forgo automatic puck recovery
  - BEFORE executing -> ESP.ready!

---
Puck Loading for Simulations
-> pucks	#returns # pucks in each tube as a Hash
-> pucks 2=>14, 3=>9	#set count of pucks for just tubes 2 & 3
-> fill!	#fill carousel tubes 2..7 with pucks, empty others
    - TubeCapacity = 22 (or 33 if using “short” pucks)
-> fill! 14, 2, 3	#fill tubes 2 & 3 with 14 pucks, empty others
- To ensure each simulation run starts from tube 2 with a full carousel:
  if ESP.simulation?
  fill!
  startTube 2
  end
  Put the above before the mission