2G ESP
Water Sampling
5/25/22 Brent Roman brent@mbari.org

---
Syringe Pump Driven Sampling
- The 2G ESP uses its 25ml Sampling Syringe
  - to pull water through filter in Collection stage
  - while monitoring syringe pressure
- Syringe stops when pressure in it falls too low
  - until pressure equalizes
  - resumes at 90% of the previous pumping rate
- Different Samplers optimized for each filter type
  - BACsampler, HABsampler, DAsampler, etc.
    - see utils/sampler.rb, utils/gensampler.rb,
    - and protocol/2G/[BAC,HAB,DA].rb

---
2G Sampler Parameters (1 of 5)
  Defined in utils/gensampler.rb:
- vacuum defaults to 10..13 [PSI]
  - target vacuum range while pulling through filter
- minRate defaults to 0.2 [ml/second]
  - Filter is clogged when filtering at minRate
    - generates vacuum > vacuum.max
- avgRate defaults to 0.2 [ml/second]
  - Constant filtering rate (used only during simulation)

---
2G Sampler Parameters (2 of 5)
  Defined in utils/sampler.rb:
- bubblePt defaults to nil [PSI]
  - Filter’s characteristic bubble point (must be spec’d)
- evacVolume defaults to 10 [ml]
  - How far to pull back SS to evacuate puck
- endVacuum defaults to 3 [PSI]
  - Vacuum must fall below this before starting next stroke
- maxDelay defaults to 40.seconds [seconds]
  - How long to wait for endVacuum to be reached
    - Filter Clogged if vacuum still > endVacuum after maxDelay

---
2G Sampler Parameters (3 of 5)
  Defined in utils/sampler.rb:
- backoff defaults to 0.9 [seconds/seconds]
  - new filtering rate after restarted due to high vacuum
- evacDelay defaults to 40.seconds [seconds]
  - how long maintain high vacuum during puck evac
- numberOfEvacs defaults to 1
  - number of times to repeat puck evacuation
- maxPumpVolume defaults to 40 [ml]
  - max vol of air pumped to reach bubblePt during evac

---
2G Sampler Parameters (4 of 5)
  Defined in utils/sampler.rb:
- equalized defaults to 1.0 [PSI]
  - Intake blocked if pressure will not equalize to within this psi
- minPSI defaults to 2.5 [PSIa]
  - Minimum sampling pressure (to avoid boiling sample)
- maxResidue defaults to 0.1 [ml]
  - don’t take another stroke if remaining vol < maxResidue
- primeVolume defaults to 7 [ml]
  - volume of initial prime stroke
- maxPumpVolume defaults to 40 [ml]
  - max vol of air pumped to reach bubblePt during evac

---
2G Sampler Parameters (5 of 5)
  Defined in utils/sampler.rb:
- meterInterval defaults to 30.seconds [seconds]
  - Flow meter reporting interval
- meterDelta defaults to .1 [Liter/min]
  - report flow changes only if > meterDelta
- exhaleConfig defaults to nil [I2C::Servo::Configuration]
  - SS configuration for exhalation

---
Sampling in the Lab
-> mySampler = Sampler.new bubblePt:26
… verify parameters are to your liking ...
-> mySampler.maxDelay=60.seconds	#increase maxDelay
-> CC.loadPuck :sh1	#load puck from startTube
-> mySampler.prime	#take the optional prime stroke
-> mySampler.sample 55	#sample up to 55ml more
-> mySampler.avgRate=2	#simulate faster!
-> mySampler.sample 125 #sample up to 125ml more
-> mySampler.puckEvac	#dry the puck
-> CC.unloadPuck	#put spent puck into discard tube

---
collectSample
-> CC.loadPuck(:sh1) {collectSample(mySampler).puckEvac}
- Loads sh1 puck from startTube
- primes and samples default goal volume with mySampler
- dries the puck
- discards it into startTube-1
- replace mySampler with any of the predefined samplers
  - BACsampler, HABsampler, DAsampler, etc.

---
Sampler Pressure Sensing
-> Sampler.intakePSI	#abs pressure above the filter
-> Sampler.outletPSI	#abs pressure below the filter
-> Sampler.PSI	#[abs pressure above, abs pressure below]
-> Sampler.calibrate	#tare pressure sensors
  - A dry puck must be clamped in the collection stage
  - Must tare after a new puck is clamped
  - For hardware debugging and initial calibration:
-> CS.status.outPress #intake pressure in raw ADC counts
-> SS.status.outPress #outlet pressure in raw ADC counts
  - Raw pressure ADC counts range 0..4095

---
External Sampling Aids
- External (Dumb) Pumps
  - Switched Lift Pumps
- Intake selector valves
  - Choose between shallow and deep sample intakes
- Passive Flow Meters
  - Verify water flowing through sample intake loop
- Active Flow Controllers
  - Combine meter with pump controller to regulate flow
  - Also controls an external intake selector valve

---
Configuring External Pumps
- Switch via ESP’s :sampler external power switch
- Sampler::PumpPrimeDelay
  - defines how long pump runs to prime external loop
  - before ESP intake opens
  - typically defined in machine’s configure.rb as:
  class Sampler
  const :PumpPrimeDelay, 3.minutes
  end

---
Intake Selector Valve
- External Valve selects between deep and shallow intakes
  - Only installed on GLERL ESPs
  - Connected as a Gripper
    - to the Sampler Dwarf microcontroller
  - Example config at end of ESPnessie’s configure.rb:
  require 'depthsel'
  :Sample.denotes DepthSelectorValve "Depth Selector Valve", :sampler, 1
  Sample.maxDuration=30 #allows up to 30 seconds for valve to move
- Valve must be explicitly positioned before sampling begun
  -> Sample.shallow	#select shallow intake
  -> Sample.deep #select deep intake

---
Passive Flow Meters
- Two types in use:
  - Analog input (by WHOI)
    - may function as a simple flow presence/absence switch
    - or uncalibrated % of full flow
    - in ESP’s Can data
  - Serial input (by NOAA NWFSC/APL)
    - on ESP’s external sampler RS-232 serial port
    - logs Liters/minute every minute
    - automatically enabled during sampling

---
Active Flow Controller
- Developed by APL with MBARI
  - combines digital flow meter with pump speed controller
  - ESP sets desired input loop flow rate during sampling
  - Powered by External Sampler power switch
  - Communicates on External Sampler serial port

---
Active Flow Controller Config
- Configuration in configure.rb:
  class Sampler
  const :PumpPrimeDelay, 3.minutes #prime before sampling
  const :PrimeFlow, 2.0..7.0 #Liters/min flow range while priming
  const :MaintainFlow, 1.0..3.5 #flow range while sampling
  end
- Configuration in corecal.rb:
  flowScale = 0.00276964 #scales counts/min flow to liters/min
  @@flow = Skew flowScale,-flowScale, format: '%.1fL/min', threshold:0,
  driver: Gauge::Flow::UWFlowControl,
  port:'/dev/extSampler', switch: :sampler

---
Active Flow Controller Use
- See lib/gauge/uwflowcontrol.rb
  -> Flow.start	#start monitoring flow
  -> Flow.selectIntake :shallow	#select shallow intake
  -> Flow.selectIntake :deep	#select deep intake
  -> Flow.intake	#verify intake selection
  :deep
  -> Flow.prime	#run pump at faster priming flow rate
  -> Flow.maintain	#run pump at slower sampling flow rate
  -> Flow.rate	#returns current flow rate in Liters/minute
  3.2
  -> Flow.disablePump	#stop pump
  -> Flow.stop	#stop pump and stop monitoring flow